## Slide 1: Industrial Gas
Measurement & Analytics
Market leading gas analyzer solutions

## Slide 2: Agenda
1. Gas analyzer solutions for industrial gas
6. Natural Gas – LNG1 / CNG2
2. Air separation & argon purification
7. Biomethane & small-scale CO2
5. Large scale CO2 – associated with H2 / syngas / ammonia
10. Summary
4. Interlinks with petrochemical industry
9. Cylinder / specialty gases
3. Hydrogen, carbon monoxide & syngas production
8. Hydrogen – emerging applications

## Slide 3: Pure class for industrial gas
Market leading gas analyzer solutions

## Slide 4: Pure class for industrial gas
Continuous Gas Analyzers
Process Gas Chromatographs
Off-Axis ICOS laser technology
Comprehensive offering for the Industrial Gases sector
Market leading gas analyzer solutions

## Slide 5: Continuous Gas Analyzers (CGA)
Most complete portfolio in the industry
September 8, 2019
Slide 5
AO2000 High-end product line
EL3000 Entry level product line
EL3060 Ex-d product line
Extractive
LS4000 / LS25 Tunable diode laser
AZ Series Zirconia O2/COe
In-situ
ACF5000 Hot/wet extractive
ACX Cold/dry extractive
Easy System Cold/dry extractive
Standard Systems
AnalyzeIT Explorer Asset management
CEM-DAS CEMS Data Acquisition
Software
Notes: High level overview
9 core measuring technologies (e.g. IR, UV, TDLAS, PMD, Electrochemical, TCD, FID, Zirconia, FTIR) with total of 18 analyzer module variations
3 product lines (extractive) – AO2000 (high-end), EL3000 (entry level), EL3060 (hazardous areas)
2 product lines (in-situ) – LS25 private labelled, LS4000 own product, AZ series for O2 and Combustibles (manufactured in UK and US)
Range of sample handling components allows offering complete standard system solutions
ACF5000 – FTIR based CEMS, clear market leader, >1700 systems
ACX – standard system for CEMS and process applications…Complete, Configurable, Consistent.
Easy System – low cost „easy“ applications, plug & play, building block concept
Asset management software and certified data acquisition for CEMS

## Slide 6: The analyzer modules
World renowned brand names built on 90 years of rich heritage
Uras
IR photometer
Original BASF license to H&B*, first usable NDIR spectrometer
CO, NO, SO2, CO2, …
Uras26
Magnos
Paramagnetic
Magnos 3 first H&B analyzer using glass dumbbell
O2 (vol%)
Magnos27/28
Fidas
Flame Ionization
Measure emissions from single-cylinder four-stroke engines
THC, TOC, VOC, CnHm
Fidas24
Limas
UV photometer
Radas 1G NDUV photometer successful launch
NO, NO2, SO2, …
Limas21/23
O2 Sensor
Electrochemical
Low initial cost option for O2 measurement
O2 (vol%)
O2 Sensor
ZO23
Zirconium Dioxide
Trace level O2 measurement with self-check feature
O2 (ppm)
ZO23
1920s
1930s
1950s
1960s
1970s
1980s
2000s
Caldos
Thermal Conductivity
First CO2 analyzer for combustion processes
H2, He, Ar, N2, …
Caldos25/27
A solid foundation for all our continuous gas analyzers
Click for more info
on analyzer modules

## Slide 7: EasyLine series (EL3000)
  Excellent price-performance ratio
  Single housing solution – rack or wall mount
  Measure up to 5 components in one unit
  Combine up to 2 sensors in one enclosure
    Uras26 NDIR
    Limas23 UV
    Magnos28 PMD O2
    Magnos27 TMD O2
    Caldos27 TCD
    Fidas24 FID
    ZO23 Zirconia O2
    Electrochem. O2 sensor
So smart, they’re simple
Uras26
ZO23
Magnos27
O2 Sensor
Magnos28
Limas23
Fidas24
Caldos27

## Slide 8: EasyLine combinations
  Uras26 + O2 sensor (up to two)*
  Uras26 + Magnos28
  Uras26 + Magnos27*
  Uras26 + Caldos27
  Limas23 + Magnos28
  Limas23 + O2 sensor
  (New!) Caldos27 + Caldos27*
  (New!) Magnos28 + Magnos28*
  (New!) Magnos28 + Caldos27*
  Fidas24 (standalone)
  ZO23 (standalone)
Up to two sensors per housing, multitude of combinations

## Slide 9: EasyLine Continuous Gas Analyzers
So smart, they’re simple
Easy configuration
Easy integration
Easy operation
Easy maintenance

## Slide 10: Easy configuration
  Easy to select the right product, tailored to your needs and stress-free set up after delivery
  Delivered tailored for each application
    Configurable measuring ranges
    Single dynamic range with two measuring ranges
    Measure flammable gases without purging
    Modular I/O based on your needs
  All commonly used functions accessible via HMI
  Every analyzer shipped with PC configuration tool for custom parameter setting

## Slide 11: Easy integration
  Good fit for each application, options to fit your needs, I/O for every eventuality
  19” rack (3HU) and wall mount housing (IP65)
  Optional pump, flow sensor and solenoid valve*
  Control up to 7 external solenoid valves
  Industry standard I/O
    0/4…20 mA for each component
    Standard 4 x DI, 4 x DO (Max. 12/12)
    Standard Modbus TCP/IP
    Optional Modbus RTU or Profibus
  Attractive solutions for hazardous applications! Click for more info…

## Slide 12: Easy operation
No complex menu structure means no training and no manual necessary
  No training/manual necessary for operation
  Intelligent 4-way navigation
  Flat menu structure
  Access to most likely functions using same button (Example: Operation Calibration Manual Cal.)
  Password protection for expert levels
OK

## Slide 13: Easy maintenance
  Innovative calibration concepts, accessible modular design, intelligent and intuitive servicing
  Optional gas-filled calibration cells reduce calibration costs by up to 95%! Click for more…
  Integrated drift checking with one year data storage (QAL3 reporting according EN14181)
  Automatic or manual calibration
  Modular design allows easy disassembly of hardware for component replacement
  Extensive self-diagnostics, clear text status messages
  Now with Dynamic QR Codes for faster fault diagnosis and repair! Click for more info…

## Slide 14: EasyLine Continuous Gas Analyzers
So smart, they’re simple
Easy configuration
  Easy to select the right product, tailored to your needs and stress-free set up after delivery
Easy integration
  Good fit for each application, options to fit your needs, I/O for every eventuality
Easy operation
  No complex menu structure means no training and no manual necessary
Easy maintenance
  Innovative calibration concepts, accessible modular design, intelligent and intuitive servicing

## Slide 15: Advance Optima series (AO2000)
Modular gas analyzer series
The AO2000 series is based on an integrated system concept
  Standardized analyzer modules can be combined to a multi-analyzer system
Fidas
Flame-ionization detector
Magnos
Paramagnetic oxygen
Limas
UV photometer
Uras
IR photometer
Caldos
Thermal conductivitiy
Notes: And the heart of all of ABB‘s continuous gas analyzers are the sensor modules:
Our flagship product the Uras is based on over 50 years of experience gained with non-dispersive infrared technology and measures up to four components simultaneously such as CO, NO, SO2 and CO2.
The Limas is our workhorse using UV measuring technology that is not affected by either water or CO2 and measures both NO and NO2 simultaneously as well as SO2 if required.
Magnos is a paramagnetic oxygen analyzer and incorporates ABB‘s revolutionary microwing technology and semi-automatic manufacturing for high precision and repeatable measurement.
Electrochemical oxygen sensors are also available if you prefer to optimize the initial cost of purchase but be aware that these will need to be replaced on average every 12-24 months.
Lastly the Fidas, using the flame ionization principle for measurement of total hydrocarbons, is world-renowned and appreciated for its robustness and reliability.
[CLICK] These sensor modules form the basis for the AO2000 series, which is an integrated system concept, and I will further explain this on the next slide…

## Slide 16: Advance Optima series (AO2000)
Modular gas analyzer series
Broadest selection of measuring principles
One central unit for up to 4 sensor modules
Measures up to 6 gas components simultaneously
Single or multi housing solution – rack and wall mount
Sensor modules up to 350m from the central unit
Sample handling components can also be integrated
Internal PLC functionality
    Valve control – autocal, stream switching
    Calculations – O2 correction, Wet to Dry
    Integrate external signals
    Cross sensitivity correction
    Program soft keys (i.e. shortcuts)
Attractive solutions for hazardous applications! Click for more info…
350 m
Module 1
Module 2
Module 3
Module 4
Module 4
Module 2
Module 3
Module 1
Notes: Up to four of these sensor modules can be combined with one central control unit measuring up to six components simultaneously.
[CLICK] Both rack and wall mount configurations are available and the sensor modules can be installed up to 350 meters away from the central unit using a digital bus architecture.
[CLICK] Completing the solution, even the ABB sample handling components can be integrated to the system via a single cable and data made available on the AO2000 display.

## Slide 17: EL3060
  One enclosure for Caldos25/27 or Magnos28
  Uras26 in separate Ex-d enclosure
  Combination of two analyzers possible
  Up to 5 measuring components
  Ex-d housing – no purging required
  ATEX 2G certified
  Click for more info about hazardous area certifications…
Compactly built
Easy to combine
Safe, compact and certified
The specialist for hazardous areas
Caldos25/27 or
Magnos28
Uras26

## Slide 18: AZ20 – Zirconia oxygen analyzer
Robust, long-life probe
  Proven cell design from over 50 years experience
  Fast response to process variations
  Stable, accurate and low drift Oxygen measurement
Easy operation
  Simple and intuitive 4-key operation
  Easy Setup menu
  Internal data-logging (up to 100 events)
Unique integrated auto-calibration system
  Integral to the probe head
  Controls the sequence and test gas flow rate
  Detects test gas availability
Robust, easy operation, integrated auto-cal feature

## Slide 19: PGC1000 – Process Gas Chromatograph
Compact
  Mounted much closer to sample point
Versatile
  Capable of a variety of vapor applications
Cost Effective
  Shelterless (if sample conditions allow)
  Low utility consumption
Quality
  Conventional analytical components
  Highly sensitive TCD* for C1-C9+, inerts and H2S
  Modularity of hardware and software
Compact, shelterless, cost effective, quality

## Slide 20: PGC1000
  Consists of three primary components
    GC module (detector, pressure sensor, column & sample loop spool, valves)
    Manifold assembly (heater plate assembly & manifold)
    Analytical processor assembly
  GC module
    2 - thermal conductivity detectors
    2 - 10 port valves (sample backflush to measure)
    1/16” micropacked or capillary GC columns
    Detector plate houses pressure sensors for EPC*
  Valve assembly
    No external tubing
    Dual 10 port valve arrangement (RCS)
    Million of cycles between failures
  Column spool assembly
Analytical module assembly

## Slide 21: PGC5000 – Process Gas Chromatograph
  Modular process GC series
    PGC5000A Controller
    PGC5000B Smart OvenTM
    PGC5000C Smart OvenTM
    Fiber optic, CANbus communications
  Full detector suite for maximum measurement capability: FID, sTCD, mTCD, FPD, DID
  Airless oven option for significant utility savings
  Superior technical solution for Total Sulfur (PGC5007B)
  Easy to use graphical, touch screen User Interface (UI)
  Industry standard communication protocols
    Native TCP/IP and Modbus TCP/IP network communications
    Native multicast type data transmission with V4 OS
State of the art, flexible, simple, powerful

## Slide 22: PGC5000
One of a kind
  The Single Board Computer (SBC), traditionally located in the PGC5000A enclosure, is now fully integrated within the analytical oven making the traditional controller and user interface (UI) optional
Significant cost savings
  Green field shelter size and analyzer cost reduction of up to 50% and 17%, respectively
  Brown field analyzer cost reduction of up to 32%
  Replaced analyzer cost reduction of up to 9%
Safety and installation flexibility
  Technicians spend more time in safe areas with remote UI flexibility
  Three UI options for the PGC5000 Integrated Controller Oven:
    Single PGC5000A for all ovens in a shelter
    PC connected via Ethernet to the analyzer network
    Class 1, Div 2 notebook PC wirelessly connected to analyzer network
Integrated Controller Oven

## Slide 23: RVP4500 – Reid Vapor Pressure analyzer
Different models to cover vapor pressure ranges
  From 0-20 psi to 0-225 psi
  Most common are 0-20 psi and 0-30 psi for shale gas
ASTM Method D5482 (off-line mode)
  RVP4500 meets the requirements of this method
Field proven reliable process analyzer
  More than 20 years of experience in refinery operations
Does not require field calibration
  Factory calibrated, minimizes support requirements
No safety purge for hazardous areas
  Minimizes consumables required in the field
Compact, field-proven, low maintenance

## Slide 24: LGR-ICOSTM 950 Series – Off-Axis ICOS laser technology
Speed
  Fast response = fast feedback loop = product on spec ($)
Sensitivity / precision
  Low LOD (ppb/ppt) = tighter control = longer catalyst life ($)
Accuracy
  True measurement = better control = product on spec ($)
Wide dynamic range
  Accurate for all process conditions = product on spec ($)
Reliability
  Stability (no drift) = no recalibration = lower maintenance ($)
  Field serviceable = minimal training = higher uptime ($)
Extraordinarily high sensitivity, fast response and over a wide dynamic range

## Slide 25: OA-ICOS: Off-Axis Integrated Cavity Output Spectroscopy
1st: TDLAS (tunable diode laser absorption spectroscopy)
[First papers mentioning tunable diode laser for spectroscopy date from the 60’s]
2nd: CRDS (cavity ringdown spectroscopy)
[First paper mentioning CRDS using TDL date from the 1988 (Anthony O’keefe, LGR founder]
3rd: OA-ICOS (off-axis integrated cavity output spectroscopy)
[1st generation OA-ICOS patent filed in 2001, published in 2004 (Los Gatos Research)]
The latest evolution in TDLAS
4th generation OA-ICOS technology in the LGR-ICOS™ analyzers: latest advancement in TDLAS
+Sensitivity
+ Accuracy and ruggedness/serviceability

## Slide 26: ABB OA-ICOS Technology
Off-axis Integrated cavity output spectroscopy – technically advanced compared to other TDLAS
Effective optical pathlength up to 25 km = sufficient absorption even at trace levels
Basic methodology
Measurement (100-1000Hz)

## Slide 27: Overview
ABB analytical solutions for Industrial Gas
Analyzer | Measuring principle(s) | Components | Units | ASU | HyCO | CNG | LNG | Bio-CH4 | CO2 | H2 | Cylinder | Speciality
EL3000 | NDIR, NDUV, PMD, TCD, FID, ZrO2, El. sensor | CO, CO2, CH4, NO, N2O, H2, Ar, O2, THC, CV/BTU | ppm to vol% | X | X | X |  | X | X | X | X | X
AO2000 |  |  |  |  |  |  |  |  |  |  |  | 
AZ20 | ZrO2 | O2 | vol% |  | X |  |  |  |  |  |  | 
PGC1000 | Gas Chromatograph – TCD | C1-C9+, CV/BTU, H2S | ppm to vol% |  | X | X | X | X | X |  |  | 
PGC5000 | Gas Chromatograph – TCD, FPD, FID, DID | C1-C9+, H2, H2S, CO, CO2 | ppm to vol% | X | X | X | X | X | X | X |  | 
RVP4500 | Reid Vapor Pressure | RVP | psi |  |  |  | X |  |  |  |  | 
LGR-ICOS 950 | Off-Axis ICOS laser technology | HCl, HF, NH3, H2S, H2O, CO, CO2, O2 | ppt to ppm | X |  | X | X |  |  |  |  | X

## Slide 28: Air separation & argon purification

## Slide 29: Air separation
1
Ambient air drawn in & compressed
Process overview
2
Precooled with water
3
Impurities removed from air
4
Main heat exchanger cools air down to approx. -175 °C
5
Through physical separation process at low temperature pure nitrogen occurs at top & pure oxygen at bottom
6
Gaseous oxygen & nitrogen via pipeline to customers
7
Liquid gases each temporarily stored in tanks
8
Transported to customers
N2
O2
Ar
Notes: Here‘s what an ASU may look like. As the name suggests, it separates air into its main component gases of oxygen, nitrogen and argon. Some Mega-ASU‘s go further and also extract the rare gases Neon, Krypton and Xenon. One or two even recover helium from the air. But, these types of ASU are really quite exotic. The vast majority of ASUs work with oxygen, nitrogen and argon.
The applications for those three gases require them to be very pure. So, the need for quality assurance gas analyzers is clear. Furthermore, within the ASU there are process-control measurement requirements. In fact, on a typical ASU and tanker-loading site you can expect to find about 20 gas analyzers. And, all of those can be manufactured and sold by ABB Analytical. ABB is very well positioned to play a leading role in the supply of gas analyzers to this sector. You have the products in the ABB portfolio that you need to succeed.

## Slide 30: Air separation unit (ASU)
Main heat exchangers
Typical process flow / measuring points
Filter
Cooler
Compressor
Local pipeline users
Air
LIN
LOX
LAR
GOX
GAN
Low pressure distillation column
Crude argon column
Pure argon column
Truck loading area/ customers
Product compression
GAR
N2
N2
N2
N2
Ar
Ar
Ar
Ar
O2 ppm
O2 ppm
O2 ppm
O2 ppm
O2 ppm
Impurities
Impurities
Impurities
Impurities
Impurities
Impurities
O2 %
O2 %
O2 %
O2 ppm
O2 ppm
Ar
Impurities
Impurities
O2 %
THC
THC
CO2
CO2
THC
O2 %
O2 %
GAR | Gaseous Argon
GOX | Gaseous Oxygen
GAN | Gaseous Nitrogen
1
2
3
5
4
8
10
9
12
15
13
14
16
17
Impurities
6
11
High pressure distillation column
HC removal
Waste nitrogen vent to atmosphere
PPU*
Molecular sieve adsorber
THC
O2 %
CO2
7
O2 ppm
LAR | Liquid Argon
LOX | Liquid Oxygen
LIN | Liquid Nitrogen
Notes: Those of you familiar with process flow diagrams will relate to this image immediately. We have transformed the visualisation of the ASU in the previous slide into the main pieces of process equipment and product flows. The typical measurement locations are marked. At each point there will often be one or more gases which require analysis for process control or quality assurance purposes.
Please note that this is a typical ASU configuration. As noted previously, some ASUs add the complexity of rare gases production. Some are simpler than the diagram shown here because they do not produce pure argon. Others do not have a molecular sieve adsorber and use other technologies for removal of carbon dioxide and moisture from the air. However, if we understand this generic ASU model we will also have a good understanding of the simpler cases.

## Slide 31: Air separation
Measuring points and ABB offering
# | Purpose | Application | Components | Ranges | Principles |  |  |  |  |  | Analyzers
1 | Safety | Air feed THC | THC* in air | 0 – 10 / 100 ppm |  | FID |  |  |  |  | Fidas24
2 | Safety | THC / CO2 impurities | THC in air CO2 in air | 0 – 10 / 100 ppm 0 – 100 ppm |  | FID |  | NDIR |  |  | Fidas24 Uras26
3 | Purity | Gaseous Oxygen (GOX) | O2 purity THC in O2 CO2 in O2 | 98 – 100 vol% 0 – 10 ppm 0 – 10 ppm |  | PMD |  | FID |  | NDIR | Magnos28 Fidas24 Uras26
4 | Purity | Gaseous Nitrogen (GAN) | O2 in N2 CO2 in N2 | 0 – 1 / 10 ppm 0 – 10 ppm |  | Zirconia |  | NDIR |  |  | ZO23 Uras26
5 | Control | Waste Nitrogen vent | O2 in N2 | 0 – 5 / 20 vol% |  | PMD |  |  |  |  | Magnos28
6 | Control / Safety | Cryogenic Distillation (N2O to avoid plugging) | O2 purity CO2 in O2 N2O in O2 THC in O2 | 98 – 100 vol% 0-10 ppm 0-100 ppm 0-500 ppm |  | PMD |  | NDIR |  | FID | Magnos28 Uras26 Fidas24
7 | Control | Nitrogen re-processing | O2 in N2 | 0 – 100 ppm |  | Zirconia |  |  |  |  | ZO23
8 | Safety | Oxygen re-processing | O2 purity | 0 – 100 vol% |  | PMD |  |  |  |  | Magnos28
Notes: Here are the first of those measurement points. We have presented them in a table and the typical analytical concentration ranges and chemical components are identified. Take a look at the terrific number of gas analysers from the ABB Analytical product range that can be sold here. Imagine landing an order to supply that suite of instruments to a new ASU build. Or, to supply a few gas analysers to a refurbishment project on an older plant. Either of those would go some way to helping you meet your annual sales targets.

## Slide 32: Air separation
Measuring points and ABB offering
# | Purpose | Application | Components | Ranges | Principles |  |  |  |  |  | Analyzers
9 | Purity | N2 purity (LIN storage) | O2 in N2 | 0 – 10 ppm |  | Zirconia |  |  |  |  | ZO23
10 | Purity | N2 purity (LIN loading) | O2 in N2 | 0 – 1 / 10 ppm |  | Zirconia |  |  |  |  | ZO23
11 | Purity | O2 purity (LOX storage) | O2 purity | 98 – 100 vol% |  | PMD |  |  |  |  | Magnos28
12 | Purity | O2 purity (LIN loading) | O2 purity | 98 – 100 vol% |  | PMD |  |  |  |  | Magnos28
13 | Control | Argon cleaning | O2 in Ar Ar in O2 | 80 – 100 vol% 0 – 20 vol% |  | PMD |  | TCD |  |  | Magnos28 Caldos27
14 | Purity | Argon cleaning | O2 in Ar Ar purity | 0 – 1 / 10 ppm 90 – 100 vol% |  | Zirconia |  | TCD |  |  | ZO23 Caldos27
15 | Purity | Ar purity (LAR storage) | O2 in Ar | 0 – 1 / 10 ppm |  | Zirconia |  |  |  |  | ZO23
16 | Purity | Ar purity (LAR loading) | O2 in Ar | 0 – 1 / 10 ppm |  | Zirconia |  |  |  |  | ZO23
17 | Purity | Ar purity (GAR compression) | O2 in Ar | 0 – 1 / 10 ppm |  | Zirconia |  |  |  |  | ZO23
Notes: And, what about these additional process streams and measurement points. Here is another list that shows the perfect match between ABB gas analyzers that you can sell and the requirements that exist on the ASU.

## Slide 33: Air separation unit (ASU)
Continuous & accurate measurements critical for efficient ASU operations
1
2
3
4
  Electricity is the critical cost input
  Over-purification consumes more power
  Accurate measurement of final product essential for efficient operation
  “Waste-nitrogen” is very high flowrate stream containing around 2 vol% oxygen
  Essential to remove CO2 and H2O from moisture traps but then vented to atmosphere
  Small percentage of oxygen in “waste-nitrogen” results in loss of valuable O2 for the ASU
  Argon is tremendously valuable by-product from ASU
  Successful (or unsuccessful) recovery can make or break economics of some ASUs
  Monitoring Argon concentration in feed to argon purification column best way to maximize recovery
Power consumption & oxygen purity
Oxygen slip
Argon slip
  Fuel-cell type oxygen analyzers are often the most hated analyzer in the instrument shack!
  Require annual replacement, easily “poisoned” by high O2, take hours to stabilize after poisoning
  During this time, ASU might have returned to stable operation and product is on spec = lost revenue!
Nitrogen purity
  Carbon dioxide and moisture would solidify at cryogenic temperatures
  Hydrocarbons from the air can explode if they accumulate to high levels in liquid oxygen
  “Break through” of these contaminants from the pre-purification stage is monitored continuously
Contaminants breakthrough
5
Notes: These pain points on a typical ASU are sensitive issues that the operator is likely to be concerned about. They may sound like no-go areas for conversation that will put everybody out of their comfort zone, but they can work in your favour. Plant managers and instrumentation engineers will be highly aware that these issues exist. Try to engage in some conversation about one or two of these topics and listen to how they describe their current situation. Try to determine their level of satisfaction with the gas analyzers they are using for these applications today. Are they really 100% satisfied? What about ease of maintenance, up-time, calibration frequency, drift, accuracy, compatibility with the latest site process control and communications systems. Is there any room for improvement in any of these areas? Hopefully you build the case for a solution to their current situation from the ABB Analytical range.
Let’s run through them.
Power consumption & oxygen purity. Electricity is the critical cost input. Over-purification of the products such as oxygen results in consuming more power than is required, however, traces of argon exist in the oxygen. So, if the ASU is configured for argon recovery (not all are) it is prudent to recover the argon as argon, so a high purity of oxygen is desirable to avoid argon wastage. So – measurement of oxygen final product purity is essential for efficient operations.
Oxygen slip. The “waste-nitrogen” stream is essential to remove CO2 and Moisture from the moisture traps on the PPU (Pre-Purification Unit) or RHE (Reversing Heat Exchanger). Waste nitrogen contains about 2% oxygen and is a very high flowrate stream which is vented to atmosphere. So, a small percentage of additional oxygen in the waste nitrogen results in loss of valuable oxygen from the ASU.
Argon slip. Argon is a tremendously valuable by-product from the ASU. Successful (or unsuccessful) recovery of argon can make or break the economics of some ASU operations. Monitoring the argon concentration in the feed from the main ASU column to the argon purification column is the best way to control argon recovery.
Nitrogen purity. The fuel-cell type oxygen analysers that are often used for ppm oxygen measurement in the pure nitrogen stream are the most hated instrument in the instrumentation shack! The cells require replacement (eg once per year) and they are quickly “poisoned” if they see a high level of oxygen. Their response time in the upward direction is quite fast... but in the downward direction after a high oxygen “poisoning” it can take a long time (hours) for the reading to stabilise. During this time, the ASU might have returned to stable operations and the product is actually on-spec... but the instrument says the product is off-spec resulting in waste. Some operators pass hydrogen into the analyser to remove the excess oxygen and return the instrument to a stable condition more quickly. But... any alternative instrument for ppm level oxygen in nitrogen will be welcomed.
Contaminants breakthrough. The Molecular Sieve Adsorber, which is also known as the PPU is designed to remove moisture, carbon dioxide and hydrocarbons from the air. CO2 and moisture would solidify in the extremely cold cryogenic temperatures of the distillation column so they must be removed to avoid process blockages. Hydrocarbons accumulate with the liquid oxygen and, as you can imagine that is not a safe mixture of chemicals. In fact, the main reason for catastrophic ASU failures in the past decades has been hydrocarbon accumulation, resulting in a number of devastating ASU explosions.

## Slide 34: Air separation unit (ASU)
Magnos28
Paramagnetic detector – 99.5 to 100 vol% O2
  Patented microwing® offers improved repeatability
  Semi-automatic manufacturing for consistent quality
  Fast response (1.3s) for improved process control
  Available in EL3000 and AO2000 series
  Jürgen Wimmer (ASU Plant Inspection, AlzChem):
  “Magnos28 is more accurate, more precise and within tighter tolerances compared to the instrument that was previously in use. We are convinced of the improved repeatability and stability of the measurement and confident that this will lead to reduced power consumption and reduced calibration related operating costs.”
Power consumption & oxygen purity
  Electricity is the critical cost input
  Over-purification consumes more power
  Accurate measurement of final product essential for efficient operation
Pain points and ABB solution

## Slide 35: Magnos28
Superior stability demonstrated over 2 month trial
Superior precision and repeatability clear to see
Greatly improved precision and repeatability for high suppressed measuring ranges

## Slide 36: Magnos28 featured in gasworld in June 2019
Jürgen Wimmer (Plant Inspection, AlzChem), who tested the Magnos28 over several weeks, commented on his satisfaction:
“We always work on continuous improvement to get the best-in-class manufacturing processes for our innovative products. The Magnos28 oxygen analyzer uses the latest technology with its unique design and manufacturing process. Therefore, it was a natural choice for us to test the Magnos28 in our air separation unit.
We established that the oxygen measurement from the Magnos28 is more accurate, more precise and within tighter tolerances compared to the instrument that was previously in use.
We are convinced of the improved repeatability and stability of the measurement and confident that this will lead to reduced power consumption and reduced calibration related operating costs on our site.”
The focal point for global industrial gas news, views and intelligence

## Slide 37: Air separation unit (ASU)
Magnos28
Paramagnetic detector – 0 to 5 vol% O2
  Patented microwing® offers improved repeatability
  Semi-automatic manufacturing for consistent quality
  Inert materials suitable for corrosive applications
  Fast response (1.3s) for improved process control
  Available in EL3000 and AO2000 series
Oxygen slip
  “Waste-nitrogen” is very high flowrate stream containing around 2 vol% oxygen
  Essential to remove CO2 and H2O from moisture traps but then vented to atmosphere
  Small percentage of oxygen in “waste-nitrogen” results in loss of valuable O2
Pain points and ABB solution

## Slide 38: Air separation unit (ASU)
Caldos27
Thermal Conductivity detector – 0 to 20 vol% Ar in O2
  Flexibility to cover wide variety of applications
  Fast response (t90 < 2s) for improved process control
  Available in EL3000 and AO2000 series
Magnos28
Paramagnetic detector – 80 to 100 vol% O2 in Ar
  Patented microwing® offers improved repeatability
  Semi-automatic manufacturing for consistent quality
  Fast response (1.3s) for improved process control
  Available in EL3000 and AO2000 series
Argon slip
  Argon is tremendously valuable by-product from ASU
  Successful (or unsuccessful) recovery can make or break economics of some ASUs
  Monitoring Argon concentration in feed to argon purification column best way to maximize recovery
Pain points and ABB solution

## Slide 39: Air separation unit (ASU)
ZO23
Zirconium dioxide detector – 0 to 1/10 ppm O2
  Trace level measurement for purity applications
  Catalytically inactive cell minimizes influence from flammable gases
  Wide dynamic range for start-up / shut-down operation
  T90 response time < 60 sec for the alternation of 2 test gases
  Automatic self-check function without test gases
  Available in EL3000 and AO2000 series
Nitrogen purity
  Fuel-cell type oxygen analyzers are often the most hated analyzer in the instrument shack!
  Require annual replacement, easily “poisoned” by high O2, take hours to stabilize after poisoning
  During this time, ASU might have returned to stable operation and product is on spec = lost revenue!
Pain points and ABB solution

## Slide 40: Air separation unit (ASU)
Fidas24
Flame Ionization detector – 0 to 10/100 ppm THC
  Excellent temp. stability through single block design
  Automatic and reliable re-ignition on flame out
  Injector for sample transport with no moving parts
  Available in EL3000 and AO2000 series
Uras26
Infrared photometer – 0 to 100 ppm CO2
  Gas-filled detectors for highly selective measurement
  Internal calibration cells minimize cost of ownership
  Available in EL3000 and AO2000 series
Contaminants breakthrough
  Carbon dioxide and moisture would solidify at cryogenic temperatures
  Hydrocarbons from the air can explode if they accumulate to high levels in liquid oxygen
  “Break through” of these contaminants from the pre-purification stage is monitored continuously
Pain points and ABB solution

## Slide 41: Hydrogen, carbon monoxide & syngas production
Often referred to as HyCO or Steam Methane Reformer (SMR)

## Slide 42: Steam Methane Reformer (SMR)
Process overview
  SMR is a process in which CH4 from natural gas is heated with steam, usually with a catalyst, to produce a mixture of CO and H2
  CH4 reacts with steam in the presence of a catalyst to produce H2, CO, and a small amount of CO2
  Steam reforming reaction is endothermic
  In a "water-gas shift reaction," CO and steam are reacted using a catalyst to produce CO2 and more H2
  In "pressure-swing adsorption" CO2 and other impurities are removed from the stream, leaving essentially pure H2
  SMR can also be used with ethanol, propane or gasoline
Pure hydrogen
Furnace
Convection section
Waste gas
Boiler
Water-gas shift reactor
Heat exchangers
Pressure swing adsorbers
Reformer tubes
Gas burners
Pipe coil
Steam added
Flue gas
Low level heat released
Sulfur removal vessel
Natural gas
Notes:   Here is a graphic representation of an SMR. In the majority of steam methane reforming processes, methane from natural gas is heated with steam in the presence of a catalyst, to produce a mixture of carbon monoxide and hydrogen. On refineries, an SMR can also be fed with other light hydrocarbons such as ethanol, propane or gasoline. The reaction is endothermic meaning that heat must be supplied to the process for the reaction to proceed. This heat is derived from the combustion of natural gas or another light hydrocarbon feedstock.
  The next stage is the "water-gas shift reaction" where the carbon monoxide and steam are reacted using a catalyst to produce carbon dioxide and more hydrogen. In a final process step called PSA, or "pressure-swing adsorption" carbon dioxide and other impurities are removed from the gas stream, leaving pure hydrogen. Similar to the case with the ASU, the gases that are being processed in the SMR are ones for which ABB Analytical has a highly suitable range of gas analyzers.
  It is increasingly the case that the carbon dioxide-rich off gas from the PSA unit is further purified and sold to applications in the food & beverage industry for freezing & chilling foods and putting bubbles into fizzy drinks. In this additional CO2 processing stage, there are even more opportunities for gas analyzer sales.

## Slide 43: Hydrogen, carbon monoxide & syngas production
Typical process flow / measuring points
Drying
station
Cold-Box
separation
column
Chilling unit package
Process-gas cooling
Chilling unit package
CO2 removal
De-sulfuration
H2- Purifier (PSA-technic) (Pressure Swing Absorption)
Natural gas ca. 90% CH4
H2
CO
Steam
H2 (1.. 2%)
O2
Synthesis gas
Steam reformer
2
1
3
4
6
5
7
8

## Slide 44: Hydrogen, carbon monoxide & syngas production
Measuring points and ABB offering
# | Purpose | Application | Components | Ranges | Principles |  |  |  |  |  | Analyzers
1 | Control | Natural Gas feed | CH4 in NG O2 in NG WI, Density, S | 0 – 100 vol% 0 – 10 vol% - |  | NDIR |  | PMD |  | GC-TCD GC-FPD | Uras26 Magnos28 PGC1000/PGC5000
2 | Emission | Emission control | NO in flue gas CO in flue gas O2 in flue gas | 0 – 250 ppm 0 – 100 ppm 0 – 10/25 vol% |  | NDIR |  | PMD |  |  | Uras26 Magnos28
3 | Control | Combustion optimization | O2 in flue gas | 0 – 10 vol% |  | Zirconia |  |  |  |  | AZ series
4 | Control | Reformer efficiency | CO in syngas CH4 in syngas | 0 – 10 vol% 0 – 50 vol% |  | NDIR |  |  |  |  | Uras26
5 | Control | Absorber inlet | CO2 in H2 | 0 – 5 / 50 ppm |  | NDIR |  |  |  |  | Uras26
6 | Control | Absorber outlet | CO2 in H2 | 0 – 5 / 50 ppm |  | NDIR |  |  |  |  | Uras26
7 | Quality | CO product quality | CH4 in CO H2 in CO CO purity | 0 – 1000 ppm 0 – 5000 ppm* Composition |  | NDIR |  | TCD |  | GC-TCD | Uras26 Caldos27 PGC5000
8 | Quality | H2 product quality | CO in H2 H2 purity | 0 – 10 ppm Composition |  | NDIR |  | GC-FID |  |  | Uras26 PGC5000

## Slide 45: Hydrogen, carbon monoxide & syngas production
Continuous and accurate measurements critical for efficient HyCO operations
  Incoming natural gas or hydrocarbon feedstock must not poison the SMR catalyst
  Essential to detect sulfur compounds and other known catalyst poisons
  Measurements highly critical for process integrity and would initiate a plant-trip to protect catalyst
  SMR relies on catalytic conversion of methane to H2 and CO
  If damaged, reaction will not take place properly
  Essential to measure products coming out of SMR to check reaction selectivity, conversion and yield
  Purification stage after SMR with Pressure Swing Adsorption (PSA) unit – semi-batch cyclical process
  Switch to new adsorber bed takes place when it is near-fully saturated with a gas (cannot adsorb)
  Monitoring “break-through” essential – indicates adsorber material aging or incomplete regeneration
Input gas quality
Selectivity, conversion & yield
PSA breakthrough
  Syngas or hydrogen final product purity must be monitored
  Excess CO2 would represent non-conformance in either case
  If hydrogen is the target (not syngas) then excess CO would also represent a non-conformance
Final product purity
1
2
3
4
Notes: These pain points on a HyCO facility are some of the most potent sales tools that you have. Plant operators and instrumentation technicians will have these issues at front of mind. If you can hook into these selling points you will demonstrate empathy with their situation, perhaps you can explore their level of satisfaction with the gas analyzers they are using for these applications and hopefully you build the case for an improved solution to their current situation from the ABB Analytical range.
Input gas quality. The incoming natural gas or hydrocarbon feedstock must be free of chemicals that might poison the SMR catalyst. Measurement to detect sulphur compounds and other known catalyst poisons is essential. These measurements would initiate a plant-trip to protect the catalyst and are highly critical for process integrity.
Reaction selectivity & yield and catalyst condition. The SMR relies on catalytic conversion of methane (or other hydrocarbons) to hydrogen and CO. If the catalyst is damaged (it deteriorates naturally over time or can be poisoned rapidly by certain chemicals) the reaction will not take place properly. So, it is essential to measure the products coming out of the reformer to check the reaction selectivity, conversion and yield.
PSA breakthrough. The purification stage after the reformer is done on a PSA (Pressure Swing Adsorpion) unit. This is a semi-batch cyclical process operation where the gases flow through an adsorber bed until that bed is near-fully saturated with a gas. If/when the bed does actually become fully saturated with the gas (not intended) then it can no longer adsorb that gas and “break-through” occurs. It is essential to monitor for breakthrough. The causes can be incomplete regeneration of the adsorber bed in the regeneration phase, perhaps due to improper temperature or poor regeneration gas purity control or it can be indicative that the adsorber material is aged and needs replacing.
Final product purity. The syngas or hydrogen final product purity must also be monitored. Excess CO2 would represent a non-conformance in either case. And, if hydrogen is the target not syngas) then excess CO would also represent a non-conformance.

## Slide 46: Air separation unit (ASU)
PGC5007B
Gas Chromatograph with FPD detector – Total Sulfur
  Liquid and vapor, from natural gas to gasoline and diesel
  Measurement in accordance with ASTM Method D7041-04
  Precision and bias data 2-5 times better than all other sulfur methods
  Only method that meets requirements for low range (0-15 ppm) diesel
  ASTM 7041-04, "Standard Test Method for Determination of Total Sulfur in Light Hydrocarbons, Motor Fuels, and Oils by Online Gas Chromatography with Flame Photometric Detection," ASTM International
Input gas quality
  Incoming natural gas or hydrocarbon feedstock must not poison the SMR catalyst
  Essential to detect sulfur compounds and other known catalyst poisons
  Measurements highly critical for process integrity and would initiate a plant-trip to protect catalyst
Pain points and ABB solution

## Slide 47: Air separation unit (ASU)
Uras26
Infrared photometer – 0 to 10 vol% CO / 0 to 50 vol% CH4
  Measure up to 4 components simultaneously
  Gas-filled detectors for highly selective measurement
  Easy to add and change measuring components in field
  Internal calibration cells minimize cost of ownership
  Available in EL3000 and AO2000 series
Selectivity, conversion & yield
  SMR relies on catalytic conversion of methane to H2 and CO
  If damaged, reaction will not take place properly
  Essential to measure products coming out of SMR to check reaction selectivity, conversion and yield
Pain points and ABB solution

## Slide 48: Air separation unit (ASU)
Uras26
Infrared photometer – 0 to 10 ppm CO
  Gas-filled detectors for highly selective measurement
  Internal calibration cells minimize cost of ownership
  Available in EL3000 and AO2000 series
PGC5000
Gas Chromatograph with TCD detector – Composition
  Modular process GC series – FID, sTCD, mTCD, FPD, DID detectors
  Airless oven option for significant utility savings
  Easy to use graphical, touch screen User Interface (UI)
  Industry standard communication protocols
PSA breakthrough
  Purification stage after SMR with Pressure Swing Adsorption (PSA) unit – semi-batch cyclical process
  Switch to new adsorber bed takes place when it is near-fully saturated with a gas (cannot adsorb)
  Monitoring “break-through” essential – indicates adsorber material aging or incomplete regeneration
Pain points and ABB solution

## Slide 49: Air separation unit (ASU)
Uras26
Infrared photometer – 0 to 10 ppm CO2 / 0 to 10 ppm CO
  Measure up to 4 components simultaneously
  Gas-filled detectors for highly selective measurement
  Internal calibration cells minimize cost of ownership
  Available in EL3000 and AO2000 series
PGC5000
Gas Chromatograph with TCD/FID detector – Composition
  Modular process GC series – FID, sTCD, mTCD, FPD, DID detectors
  Airless oven option for significant utility savings
  Easy to use graphical, touch screen User Interface (UI)
  Industry standard communication protocols
Final product quality
  Syngas or hydrogen final product purity must be monitored
  Excess CO2 would represent non-conformance in either case
  If hydrogen is the target (not syngas) then excess CO would also represent a non-conformance
Pain points and ABB solution

## Slide 50: Interlinks with petrochemical industry

## Slide 51: Methanol production
ASU for Oxygen and SMR for Hydrogen
Oxygen
Desulfurization
Pre-reformer
Autothermal reformer
Methanol synthesis
Natural gas
Methanol distillation
Pure methanol
Air separation
Hydrogen recovery unit
Air
Process Steam

## Slide 52: Ammonia production
SMR for Hydrogen & ASU for Nitrogen (Oxygen can be used in other processes)
PSA
Hydrogen compression
PSA
Ammonia synthesis
HP H2
NH3
LP H2
ASU
N2 (from ASU or pipeline)

## Slide 53: Ethylene Oxide production (feed to MEG, PET/Polyester etc.)
ASU for Oxygen
Primary reactor
Reactor
Secondary reactor
Stripping column
Scrubber
Scrubber
Scrubber
Desorption column
Off-gas purification
Air purification
Stripping column
Distillation column
Ethylene
Inhibitor
Vent gas
Air
Ethylene
Inhibitor
Oxygen
Vent gas
CO2
Ethylene oxid
Oxygen process
Air process
Ethylene oxide
Light products
Ethylene oxide & water
Ethylene glycol

## Slide 54: FCC catalyst regeneration
ASU for Oxygen
Flue gas to co boiler
Spent catalyst from reactor
Regenerator standpipe
Regenerated catalyst to reactor-riser
+30 in. H2O pressure
Air blower
Oxygen injection valve
Vaporisation unit
Oxygen analyzer probe
Fluidized catalyst bed
Air grid
Dip leg
2 stage cyclone separators
Plenum chamber
Liquid oxygen tank
Basis schematic of FCC plant
FCC-reactor
Boiler
Regenerator
Fractionation
Recycle oil
Vacuum gas oil
Oxygen
Regeneration air
Steam
Boiler feed water
Redidual oil
Cracked gas oil
Cracked gases
Cracked gasoline
Steam
Catalyst. loaded
Catalyst. regenerated
Notes: During the trip up riser, the cracking catalyst is “spent” by reactions which deposit coke onto the catalyst and greatly reduce its activity and selectivity. The “spent” catalyst is disengaged from the cracked hydrocarbon vapors and routed to a stripper where it is contacted with steam to remove hydrocarbon remaining in the catalyst pores. The “spent” catalyst then flows into a fluidized-bed regenerator where air (or in some cases air plus oxygen) is used to burn off the coke in order to restore catalyst activity and provide the necessary heat for the next reaction cycle, cracking being an endothermic reaction. The hot “regenerated” catalyst then flows to the base of the riser, repeating the cycle.

## Slide 55: Refinery liquid fuels desulphurisation
Hydrogen for hydrodesulphurisation / hydrotreating
Stripper tower
Hydrotreating reactor
Amine treatment unit
Separator
Furnace
Hydrocarbon feed H2
Gas, H2S, water vapor
Desulfurized hydrocarbon product
H2S
Make-up H2
Recycle H2

## Slide 56: Coal or refinery coke gasification
ASU for Oxygen
CO
H2
Steam
Coal
Air separation unit
Coal gasification
Syngas cooling and heat recovery
Water gas shift
Particulate filter
Acid gas removal
H2 and CO separation

## Slide 57: Gasification of coal and refinery residue
ASU and HyCO plants
Coal supply
PLAR
Sulphuric acid off-take
CO2 to EOR
Power off-take
Urea off-take
SG
H2S
GAN
GAN
NH3
CO2
H2
GOX
CO2
H2
GAN
Shift, LTGC, Rectisol, NWU
Ammonia loop
Power plant
Balance of plant (U&O)
Coal storage & preparation
Gasification
ASU
H2SO4
Urea plant
CO2 compression

## Slide 58: Integrated Gasification Combined Cycle (IGCC) power generation
ASU Oxygen
Solids
Gasifier
Gaseous constitutes
Electric power
Steam turbine
Generator
Gas turbine
Generator
Electric power
Air separator
Shift reactor
Synthesis gas conversion
Solid by-product
Solid by-product
Sulfur by-product
Coal, petroleum coke, biomass, waste, etc.
Fuel cells
Fuels & chemicals
Stack
Hydrogen separation
Heat recovery steam generator
Combustor
Particulate removal
Gas cleanup
Electric power
Air
Air
Steam
Steam
Air
Compressed air
Oxygen
Hydrogen
Carbon dioxide sequestration

## Slide 59: Other applications
Petrochemical industry
Flue gas oxygen for efficiency
  O2 and CO (optional)
Determine regenerator efficiency
  O2, CO, CO2
Regulatory requirement
  CO, CO2, NO, NO2, SO2, TOC
Process heaters / Furnaces
Catalyst regeneration
Emission monitoring
  H2/HC stream and C1 (optional)
Combustion efficiency, DeNOx control, emission monitoring
  O2 and CO (optional)
  CO, CO2, NO, NO2, SO2, TOC, NH3
Flare lines, vent headers, vapor recovery, leak detection
  O2
  TOC
Hydrogen recycling
Power plants
Safety issues

## Slide 60: Large scale CO2
Associated with H2 / syngas / ammonia and breweries

## Slide 61: CO2 as by-product from large SMRs
Linked to refineries and ammonia production
Feed
Steam
Ammonia product
Carbon dioxide removal
Methanation
Steam-methane catalytic reformer
Sulfur removal
Shift conversion
Ammonia synthesis

## Slide 62: CO2 from fermentation process
Often used in captive cycle within brewery or at bottling plant
Fermentation
Gas balloon
Scrubber
CO2 condenser
CO2 compressor
CO2 storage tank
Stripper reboiler (optional)
Refrigeration unit
CO2 evaporator
Foam trap
CO2 to consumers
Foam drain
Purge
Activated carbon filter/ driers
Water
Water

## Slide 63: Large scale CO2 production
Solution
  Uras26 NDIR photometer (90…100 vol% CO2)
Solution
  Uras26 NDIR photometer (CO / CH4)
  Limas21/23 NDUV photometer (NO / NO2)
  Fidas24 flame ionization detector (THC)
  Magnos28 paramagnetic detector (O2)
CO2 purity
  Carbon Dioxide (CO2)
CO2 impurities
  Carbon monoxide (CO)
  Methane (CH4)
  Nitrogen Oxide (NO)
  Nitrogen Dioxide (NO2)
  Total Hydrocarbons (THC)
  Oxygen (O2)
Gas analyzer solutions

## Slide 64: Natural gas
Compressed (CNG) and Liquefied (LNG)

## Slide 65: Liquefied Natural Gas (LNG) plant
Process flow diagram
Natural gas
Acid gas treatment
LNG
LPG
Condensate
Liquefaction
Gas wells
Reception
Condensate stabilization
Acid gas removal
Fractionation
Dehydration/ Mercury removal
Storage and loading
Utilities
Propane heat exchanger
Carbon filter
Dehydrator
Scrubber
Dry gas filter
Propane chiller
Amine treater
Inlet scrubber
Methane compressor
Air fin heat exchanger
Ethylene compressor
Air fin heat exchanger
Propane compressor
Air fin heat exchanger
Storage tanks & loading pumps
Raw gas
Vapor from ship when loading
To leading facility
Propane surge
Ethylene surge
Inlet meter
NGL/C2+
Ethylene heat exchanger
Methane heat exchanger
Plant fuel

## Slide 66: ABB solutions for LNG plants
  ABB offers analyzers for a wide range of applications within an LNG Plant
  The advantage of our analyzers are a small footprint when using our integrated controller version and low utility consumption by using Airless oven technology
  PGC5000 Process Gas chromatograpph is used to analyze the inlet feed gas through all stages of the process
  In the Sulfur removal unit (Acid Gas removal) the fractionation process with the separation of Propane, Butane, Pentane, the aromatics (BTEX) and the removal of C6+ until the final quality control of the LNG
  RVP4500 RVP analyzer is used to measure the vapor pressure of the condensate
Trust ABB to solve your toughest measurement challenges

## Slide 67: LNG Plant
Typical measuring points and ABB offering
# | Purpose | Application | Components |  | Analyzers
1 | Control | Raw gas inlet | BTU, Density, N2, CO2, Water, CH4, C2H6, C3H8, C4H10, C5H12, C6+ C7 to C10, Aromatics | % to ppm level 20 to 2500 ppm ranges | PGC5000
2 | Control | Acid gas removal outlet | H2S CO2 | 0 – 50 ppm 0 – 500 ppm | PGC5000
3 | Control | Methane Cold box outlet Aromatics Removal | N2 Benzene Toluene, Ethylbenzene etc | 0 – 10 % ppm Level | PGC5000
4 | Control | C6+ Removal in different factions,(Ethen, Propane, Butane) | C6+ | 0 to 100 ppm | PGC5000
5 | Control | C1 – C5 Heavy reflux | CH4 C6+ | 0 – 50 vol% 0-200 ppm | PGC5000
6 | Control | De Butanizer | Butane Propane | 0 - 70 % 0 - 70% | PGC5000
7 | Control | Vapor Pressure in Condensate | Vapor Pressure | 0 – 2,5 bar (0 – 20 PSI) | RVP4500

## Slide 68: Biomethane & small-scale CO2

## Slide 69: Biogas purification
Sophisticated analytical needs due to final product quality
Biogas reactor
Gas cleaning
Bio-reforming
Methanol synthesis
De-gassed manure
Manure & straw
Electrolysis
Methanol
H2O
CO2
CH4
H2S
CO2
CH4
CO
H2
H2
S
Destillation column
Storage tank
CO2 scrubber
Dehydrator
COS filter
H2S
filter
Water scrubber
Booster blower
CO2 compressor
Refrigerant
outlet/ inlet

## Slide 70: Biogas purification
Solution
  Uras26 NDIR photometer (CO2 / CO / CH4)
  Limas21 NDUV photometer (H2S)
  Fidas24 flame ionization detector (THC)
  Caldos27 thermal conductivity detector (H2)
  Magnos28 paramagnetic detector (O2)
Alternative solution
  PGC1000 process gas chromatograph (All)
  PGC5000 process gas chromatograph (All)
Components
  Carbon dioxide (CO2)
  Carbon monoxide (CO)
  Methane (CH4)
  Hydrogen Sulfide (H2S)
  Total Hydrocarbons (THC)
  Hydrogen (H2)
  Oxygen (O2)
Gas analyzer solutions

## Slide 71: Developing Market: CO2 recovery plants
Typical process flow
CO2 storage tank
Cooling tower
Drier
CO2 compressor
Scrubber tower
Absorber tower
Gas cooler
Condensate separator
Scrubber water system
Purifier
Cooler
CO2 condenser
Fired reboiler
Stripper tower
Carbon filter
Heat exchanger
Fuel air
Exit gas
Caustic tank & dosing
Refrigerant condenser

## Slide 72: Small scale CO2 production
Solution
  Uras26 NDIR photometer (90…100 vol% CO2)
Solution
  Uras26 NDIR photometer (CO / CH4)
  Limas21/23 NDUV photometer (NO / NO2)
  Fidas24 flame ionization detector (THC)
  Magnos28 paramagnetic detector (O2)
CO2 purity
  Carbon Dioxide (CO2)
CO2 impurities
  Carbon monoxide (CO)
  Methane (CH4)
  Nitrogen Oxide (NO)
  Nitrogen Dioxide (NO2)
  Total Hydrocarbons (THC)
  Oxygen (O2)
Gas analyzer solutions

## Slide 73: Hydrogen
Emerging applications

## Slide 74: Hydrogen
  Germany plans to build 400 vehicle fueling stations
  Industrial gas majors Westfalen, Linde and Air Liquide involved in clean energy partnership (CEP) which includes energy majors Shell and OMV with goal of achieving large-scale transition to H2 as vehicle fuel
  Many smaller niche and start-up companies such as McPhy and Hydrogenics also involved
  Major automotive manufacturers such as Hyundai, Toyota, BMW and Daimler developing H2-fueled cars
  Rigorous quality standards related to Hydrogen
  Clearly, analytical instrumentation will be required to validate required quality levels
Emerging applications – electrolysis, transport fuel, gas-grid injection
Hydrogen increasingly being used beyond conventional applications
  Project to be completed in 2021
  Re-use existing natural gas grid
  For homes and businesses to use for heating & cooking
  Pilot to replicated throughout UK
  Highly likely new analytical instrumentation required
UK ‘H21 Leeds Citygate’ project to generate H2 that will be compressed into gas distribution grid
  France to inject H2 at 2-6% concentrations to gas grid
  Netherlands to use H2 to fire gas turbine power station
  H2 produced using modern electrolysis technologies
  Also possible to crack ammonia to produce H2
Other developments

## Slide 75: Hydrogen production
Solution
  Caldos27 thermal conductivity detector (H2)
Solution
  Uras26 NDIR photometer (CO / CH4)
  Fidas24 flame ionization detector (THC)
  Magnos28 paramagnetic detector (O2)
Alternative solution
  PGC1000 process gas chromatograph (All)
  PGC5000 process gas chromatograph (All)
H2 purity
  Hydrogen (H2)
H2 impurities
  Carbon monoxide (CO)
  Methane (CH4)
  Total Hydrocarbons (THC)
  Oxygen (O2)
Gas analyzer solutions

## Slide 76: Cylinder / specialty gases

## Slide 77: Cylinder gases
ABB offers analytical solutions for all types of filling station
  Electronic Special Gases
  Specialty Gases & Certified Mixtures
  All products in group 3, plus…
  Medical nitrous oxide
  High purity industrial or specialty gases
  Medical & industrial gas mixtures
  Medical oxygen
  LPG / Propane
  Pure industrial gases
  Some very common industrial gas mixtures
Regional or national hubs
Medium number of sites with some medium complexity instruments. Multiple sites in a large country.
Local production sites
Many basic cylinder filling sites close to their regional markets. Each having few basic analyzers.
Global, regional or national hubs
Few highly sophisticated sites with many high value analyzers, can export products internationally.
Filling stations hierarchy
1
2
3
Notes: The complexity cylinder filling stations can be thought of as a hierarchy.
Global, regional or national hubs for specialty gases and electronic specialty gases.
Regional or national hubs for industrial or medical gases.
Local production sites for industrial gases.
There are lots of smaller simpler sites. This is because they produce low value products that do not have sufficient value to travel long distances. Also, the barriers to entry for filling simple products are low, so many smaller competitors enter the market with local cylinder filling stations. The QA / QC requirements will be basic but a few analysers are likely to be used for raw material quality assurance and to check simple gas mixture compositions in the percentage range, for example 20% CO2 in Argon.
The middle layer of the pyramid is mid complexity sites. There are fewer of them than the simplest sites and they make higher value products with a higher level of complexity that have a higher value and can profitably travel longer distances. A broader range of instrumentation with some interesting analytical requirements will be common.
At the peak of the pyramid are a small number of highly sophisticated sites. These are the specialty gases mixtures labs and the electronic special gases facilities. You might find as many instruments in the small number of these sites as all of the mid-scale facilities put together. However the requirements are very specialized and the end-users are highly demanding. The simple and mid complexity sites are more likely to specify standardized equipment or to have their gas analyzers specified by a central engineering team. The complex sites are likely to engage in this process for themselves and treat each analytical requirement as unique.

## Slide 78: Medical gases
Analytical techniques tightly defined by national and international pharmacopeia standards
HiTech Engineered Solutions is one manufacturerer of nitrous oxide production plants
O2 is generally manufactured as bulk liquefied gas on ASU plants and delivered to filling plant
N2O is produced by thermal decomposition of ammonium nitrate
Two main gases are oxygen and nitrous oxide
Gas mixture must be analysed to ensure true 50/50
Either supplied as pure gases or 50% N2O in O2
Main gas must also be confirmed to be the stated gas (i.e. O2 analyzed to ensure not N2 or another gas)
Specification of analyzers also defined very tightly

## Slide 79: Industrial / welding gases / gas mixtures
Filling plants built and operated by industrial gases companies, not by third parties
Composition of the gas mixture must be analyzed to ensure it meets required blend composition specification within the acceptable tolerances
Industrial cylinder gases are either supplied as pure gases or as gas mixtures, e.g. gas mixtures of argon and carbon dioxide are very common for welding
Main gases used for industrial applications are oxygen, acetylene, argon, nitrogen and carbon dioxide
Carbon monoxide, helium and hydrogen are also used
Sources for most of these gases are bulk liquefied gases from ASU’s or bulk liquefied carbon dioxide deliveries from a local CO2 source
Acetylene is produced at specialist filling facilities from reaction between water and calcium carbide

## Slide 80: Filling stations
Common analyzer requirements for industrial gas filling stations
General purpose area | 
Analyzer | Measuring Range
AO2000 Caldos27 N2/He | 0 - 100 Vol%
AO2000 Caldos27 Ar/He | 0 - 100 Vol%
AO2000 Caldos27 N2/Ar | 0 -100 Vol%
Hazardous area (ATEX Kat 2G) | 
Analyzer | Measuring Range
EL3060 Caldos27 H2/Ar | 0 - 50 Vol%
EL3060 Caldos27 CH4/Ar | 0 - 20 Vol%
EL3060 Caldos27 C2H4/N2 | 0 - 16 Vol%
EL3060 Caldos27 H2/He | 15 - 50 Vol%
EL3060 Caldos27 H2/Ar | 0 - 30 Vol%
EL3060 Caldos27 H2/N2 | 0 - 100 Vol%
EL3060 Caldos27 H2/Ar | 0 - 100 Vol%

## Slide 81: Food and beverage gas mixtures
Product group that sits at the edge between industrial gases and specialty gases
Gas chromatography often used to detect and quantify range of organic and inorganic impurities in the CO2
On the other hand, food application and related quality control aspects means handled like “specialty” gases
Due to production pathway of CO2, there are additional quality control precautions associated with that gas
Blends for oxygen, nitrogen and carbon dioxide
In some countries, carbon monoxide is added to some red meat packaging gas mixtures at low % level concentrations
Fruit ripening gas is a 4% ethylene in nitrogen gas mixture
Gas components used, high production volumes and the % level ranges are like “industrial” cylinder gas mixtures

## Slide 82: Specialty gases / gas mixtures
Sold with a certificate of analysis which states the precise composition of the gas
Some extreme niche mixtures (e.g. TO14, TO17 VOCs or elemental Hg) produced in 1 or 2 locations globally
More complex specialty gas mixtures produced at hub locations on a regional basis (e.g. Western Europe, Eastern Europe, South-East Asia)
Number of instruments depends on complexity of gases
Typically more than 20 analyzers per filling operation
Range from simple NDIR and Paramagnetic analyzers for CO2, CO and O2 to more sophisticated GC-FID setups for complex multi-component hydrocarbon blends
Most industrial gas companies produce some simple specialty gas mixtures locally in most countries

## Slide 83: Electronic specialty gases – bulk and cylinder
Ultra high purity gases require higher accuracy niche analytical instrumentation products
Emissions must also be controlled and monitored
For example, cavity-ring-down-spectroscopy (CRDS) or LGR Off-Axis ICOS technology
Beyond high purity air gases, chemical gases such as silane, ammonia, hydrogen bromide, trichlorosilane are used – highly specialized for silicon chip manufacturing
Industrial grade N2 might be 99.98% or N3.8 purity
Additional purification stages required or dedicated ASUs are built – more expensive to construct/operate
Electronics grade might be 99.9999% or N6.0 purity
Higher accuracy instrumentation with very low detection limits (ppb/ppt) are usually required

## Slide 84: Summary

## Slide 85: Conclusion
Pure class for industrial gas
ABB offers comprehensive gas analyzer solutions for wide range of industrial gas applications
Partnering together brings opportunities to improve profitability and lower environmental impact

## Slide 87: Uras26
Dual-beam, Luft-type NDIR (2.5 – 8 μm)
Capabilities
  Components: CO, NO, N2O, SO2, CO2, CH4, …
  Ranges*: 0 – 100 ppm up to 100 vol%
Advantages
  Measure up to 4 components simultaneously
  Measure up to 2 streams continuously
  Gas-filled detectors for highly selective measurement
  Easy to add and change measuring components in field
  Internal calibration cells minimize cost of ownership
Non-dispersive infrared (NDIR) photometer
◄
Emissions
Safety
HazLoc

## Slide 88: Limas23
Gas filter correlation UV (200 – 600 nm)
Capabilities
  Components: NO, NO2, SO2
  Ranges*: 0 – 50 ppm up to 100 vol%
Advantages
  Measure up to 3 components simultaneously
  Insensitive to H2O and CO2 interference
  Simultaneous NO and NO2 measurement w/o converter
  Internal calibration cells minimize cost of ownership
  Four beam detection simplifies failure diagnosis
Non-dispersive ultra violet (NDUV) photometer
◄
Emissions
Safety

## Slide 89: Magnos28
Paramagnetic detector with microwing®
Capabilities
  Components: O2
  Ranges*: 0 – 0.5 vol% up to 100 vol%
Advantages
  Patented microwing® offers improved repeatability
  Semi-automatic manufacturing for consistent quality
  Inert materials suitable for corrosive applications
  Fast response (1.3s) for improved process control
  Suppressed ranges (99.5-100 vol%) for purity applications
  SIL declaration - EN 61508 (2010) part 2 - click for more…
Paramagnetic detector (PMD)
◄
Emissions
Safety
HazLoc

## Slide 90: Magnos27
Heavy-duty thermomagnetic detector
Capabilities
  Components: O2
  Ranges*: 0 – 10 vol% up to 100 vol%
Advantages
  Reliable measurement for flue gas applications
  Robust cell design ideal for corrosive samples
  Easy to clean even after liquid/acid contamination
  Very popular for cement and metals applications
Thermomagnetic detector (TMD)
◄
Safety

## Slide 91: Caldos27
Fast thermal conductivity detector
Capabilities
  Components: H2, He, Ar, N2, …
  Ranges*: 0 – 1 vol% up to 100 vol%
Advantages
  Flexibility to cover wide variety of applications
  Fast response (t90 < 2s) for improved process control
  Standard application for H2 cooled turbogenerator
  Standard application for blast furnace with Uras26
Thermal conductivity detector (TCD)
◄
Safety
HazLoc

## Slide 92: Fidas24
Flame ionization detector
Capabilities
  Components: THC, TOC, VOC, CnHm
  Ranges*: 0 – 10 ppm up to 10,000 ppm
Advantages
  Excellent temp. stability through single block design
  Automatic and reliable re-ignition on flame out
  Injector for sample transport with no moving parts
  Optional heated connector block to avoid cold spots
  Optional internal pump for mobile applications
Flame ionization detector (FID)
◄
Emissions
Safety

## Slide 93: O2 sensor
Electrochemical sensor
Capabilities
  Components: O2
  Ranges*: 0 – 10 vol% up to 25 vol%
Advantages
  Low initial cost option for O2 measurement
  Combined as add-on with Uras or Limas modules
  Up to two sensors possible together with Uras**
  6 - 24 months lifetime depending on gas composition
  Certified for continuous emission measurement
Electrochemical sensor
◄
Emissions
Safety

## Slide 94: ZO23
Zirconium dioxide sensor
Capabilities
  Components: O2
  Ranges*: 0 – 1 ppm up to 250,000 ppm
Advantages
  Trace level measurement for purity applications
  Catalytically inactive cell minimizes influence from flammable gases
  Wide dynamic range for start-up / shut-down operation
  T90 response time <60 sec for alternation of 2 test gases
  Automatic self-check function without test gases
Zirconium dioxide detector
◄
Safety
Back to previous section

## Slide 95: Lower your cost of ownership
Up to 95% reduction in calibration costs
Notes: Talking of cost of ownership, I would now like to dig deeper into this topic and show you how you can reduce your operating costs by as much as 60%...

## Slide 96: Calibration concepts
  Calibration is necessary for all gas analyzers
    Some require more frequently than others
  Effort/costs vary significantly depending on approach
  Regulatory aspects can also have a big impact
    e.g. daily validation required according to US EPA
  Costs of gas cylinders should not be underestimated
Worth paying attention to the details
◄
Notes: Calibration is necessary for all types of gas analyzers although some require more frequent calibration than others.
[CLICK] The effort and costs can vary significantly based on the approach taken towards calibration.
[CLICK] Add in the regulatory aspect and it may be necessary to validate the CEMS even more frequently, especially where the US EPA requirements are followed because here daily validation is mandatory.
[CLICK] The cost of gas cylinders should not be underestimated and there are significant cost savings to be achieved with the right approach.

## Slide 97: Calibration concepts
Test
gas 1
Calibration cells minimize operating costs and handling of test gases
Save money by picking the right solution
Flue
gas
Ambient air
Analyzer
Cooler
Pump
Calibration cells
Automated
Manual
Automated with Calibration Cells
Test
gas 2
Test
gas 1
Test
gas 2
Flue
gas
Ambient air
Analyzer
Cooler
Pump
Flue
gas
Analyzer
Cooler
◄
Notes: The most basic approach, which is very labor intensive and costly, is to perform manual calibration where the test gases are brought in and connected up to a valve in the system for introduction to the analyzers. The one benefit of this approach is that test gases can be shared between multiple systems.
[CLICK] This process can be automated through the use of solenoid valves and programmable logic but this adds additional hardware and means the test gases can no longer be easily shared between multiple systems.
[CLICK] The most cost effective approach is to automate the process using internal calibration cells for the span point and ambient air for the zero point, thereby eliminating manual work, unnecessary hardware and the high and ongoing cost of test gas cylinders.

## Slide 98: Calibration cells
Optionally available for AO2000, EL3000, EL3060 and ACX
Up to 95% reduction in calibration costs!
Suitability
Manufacturing
Proven technology
  Gas-filled cell
  Tightness guaranteed by proprietary sputter and soldering technique
  30 years’ manufacturing experience
  Alternative to flowing test gas
  Available in both Uras and Limas photometers
  Compliant with EN 14181 and US EPA 40 CFR Part 60
  Tested by TÜV for over 10 years
  Stability superior to many test gases
  Drift < 0.5% per year
  Certified acc. EN 15267
◄
Notes: Optionally available in the AO2000 series, EL3000 series and also the ACX system solution, internal gas-filled calibration cells have proven to be extremely popular with our customers.
Glass windows are essentially welded to the gold-plated, metal housing using a proprietary sputter and soldering technique, thereby guaranteeing that the gas cannot escape.
A real test gas with known concentration is connected through the tubes on either end and then sealed inside the cell and ABB has over 30 years of experience with this manufacturing process.
[CLICK] Calibration cells are an acceptable alternative to flowing test gas for the vast majority of emission monitoring regulations including the European EN14181 and US EPA 40 CFR Part 60.
[CLICK] Two cells were tested side-by-side by TÜV in Germany for over 10 years and they found that the stability was superior to many cylinder test gases with drift of less than 0.5% per year.
The calibration cells are specifically mentioned in the AO2000 and EL3000 certifications according to EN15267.
[CLICK] We’ll see later on that utilizing internal cells can reduce calibration costs by as much as 95%.

## Slide 99: Return on investment (ROI) calculation
Calibration cells
Ask your local sales representative for this tool and see how much you could save!
◄
Back to previous section
Notes: To see how much you could save, please contact your local sales representative who can take you through a return on investment calculation based on your individual situation.

## Slide 100: Advanced digital services
Dynamic QR Codes + myIB + myABB
Notes: The term digital is often over-used and the common misconception is that there is little that is really useful behind it, but I‘d like to introduce you to some of the latest innovations by ABB that can be used today and provide real value in how you interact with and maintain your assets…

## Slide 101: Common scenario
Communication ‘ping-pong’
Our analyzer is not working. Can you please help us?
It’s a Uras analyzer. Where can I find additional information?
I cannot access the type plate since the unit is in a cabinet…
We would be glad to. Can you please provide some more information?
The serial number is available on the type plate of the analyzer.
◄
Notes: Let‘s start with a common scenario that we call communication ‘ping-pong’…
Imagine you have an analyzer on your plant and it is showing an alarm and you need support to identify what the problem is.
You call your supplier and the following example unfolds:
[CLICK] Our analyzer is not working, can you help us?
[CLICK] Of course, we‘d be glad to. Can you provide some additional information about the equipment?
[CLICK] Sure, it‘s a Uras analyzer. What else do you want to know?
[CLICK] Ok, please give me the serial number on the type plate.
[CLICK] Hmm…that‘s difficult to find as the analyzer is installed in a rack. Is there another way?
You can easily consume a lot of time just trying to obtain basic information about the device before you can even begin to start diagnosing what the problem is.

## Slide 102: Dynamic QR Codes
  Standard QR Code reader applications
  ABB application “my Installed Base” (myIB)
  Static information for system identification
  Dynamic information on system configuration and analyzer health status
What are Dynamic QR Codes?
Integration of dynamically generated QR Codes on analyzer display
Contained information
Compatible with mobile devices
Introduction
◄
Notes: You are no doubt familiar with QR Codes as we see them everywhere these days but dynamically generated QR codes embedded on the display of the analyzer is something new.
[CLICK] These QR codes are automatically generated by the device based on the current status of the analyzer and not only contain static information like the serial number, fabrication date and software version, but also dynamic or changing information such as current measured values, temperature, pressure, flow and drift values.
[CLICK] You can simply take a photo of the QR code and send it to ABB, use any standard code reader or the ABB my Installed Base app and in a matter of seconds you can provide, in one communication, all the relevant information about the device to your ABB service department enabling them to immediately start advising you how to go about fixing the problem.

## Slide 103: Dynamic QR Codes
High simplicity: standardized, platform-independent
Increased availability: case-specific information and individual support
Improved installed base management
Benefits
1
2
3
Faster fault reporting, diagnosis and repair = increased availability
◄
Notes: In summary, Dynamic QR Codes are simple to use and platform independent, provide case specific information which allows for individual support and, as you will see in the next few slides, also improved installed base management.
[CLICK] Faster fault reporting allows us to diagnose the failure and suggest corrective actions much quicker, which results in a faster repair and ultimately increased availability.

## Slide 104: myIB
Mobile installed base management application
Report a problem
Request for service*
Register your device
◄
Notes: myIB is ABB‘s mobile installed base management application available in both the App Store and Google Play.
[CLICK] It’s free to download and allows you to register your device using the Dynamic QR Codes or keying in the serial number, report a problem and even request for a service engineer to visit your site.

## Slide 105: You can easily download application from stores
Download my Installed Base app from App Store and Google Play
ABB Information System AG
◄

## Slide 106: myABB
24/7 access to ABB
Find the right information, saving time and money
Optimize lifecycle management
Minimize capital outlays and optimize operational budgets
Improve assets performance
Plan maintenance operations and minimize downtime
Boost operational efficiency
Make the right operational decisions
Your gateway to services
Single online entry point to services based on your needs
Click for video introduction
Notes: myABB is an online portal that acts as your gateway to services from ABB.
[CLICK] It’s available 24/7 and enables you to customize the view to only see what is relevant for you.
Through myABB you can:
View the lifecycle status of all your devices providing a complete overview of your installed base
Download technical documentation such as analyzer data sheets, service reports, standard documentation and system drawings
Get an overview of completed or planned service events and receive specific recommendations for your devices
You can view spare parts information and pictures and in a future development also order parts online
For the sake of time we will not show the video now but I encourage you to look and familiarize yourself with the portal.

## Slide 107: myABB
Example: Global cement producer
◄
Notes: Here is an example for one major cement producer showing:
[CLICK] Summary of service events going back to 2014
[CLICK] A map of all their sites with the possibility to click on each flag to view more information
[CLICK] A place to search for information and register products
[CLICK] A summary of previous service activities and the device these are related to
It‘s a very powerful tool and the capabilities will continue to grow over time.

## Slide 108: Advanced digital services
Faster diagnosis, faster repair, higher availability, lower cost of ownership
Dynamic QR Codes
myIB app
myABB portal
◄
Back to previous section
Notes: Putting these together, Dynamic QR Codes combined with the myIB app and myABB customer portal offer you faster fault diagnosis, faster repair, higher availability and lower cost of ownership.

## Slide 109: Hazardous applications
Your safety is our priority

## Slide 110: Hazardous applications
CSA Class I, Div. 2 |  |  |  |  |  |  |  |  |  |  |  |  | * |  |  |  | 
EAC Zone 1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
IECEx / TIIS / NEPSI / KC‘s Zone 1 |  |  |  |  |  |  |  |  |  |  |  |  | + |  |  |  | 
 | Uras26 | Magnos206 | Magnos28 | Caldos25 | Caldos27 | Uras26 | Limas21 | Magnos206 | Magnos27 | Magnos28 | Caldos25 | Caldos27 | LS4060 | Uras26 | Magnos206 | Caldos27 | Magnos28
Zone 1 Category 2G |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Zone 2 Category 3G Flammable Gases |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Zone 2 Category 3G |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Ex free Flammable Gases |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Your safety is our priority
AO2000
EL3060
EL3000
LS4060
◄

## Slide 111: General Purpose
Flammable Gases
  Manufacturer Declaration (tested and approved by TÜV)
  EL3000 analyzer modules
  Uras26
  Caldos27
  Magnos28
  Stainless steel tubing (Uras) and direct access analyzers (Caldos/Magnos)
  IP20 Housing (EL3020), air convection
  IP65 Housing(EL3040), housing purge necessary
EL3000
Zone 2
Zone 1
Zone 0
Purged
◄

## Slide 112: Zone 2
  BVS 16 ATEX E 085 X (II 3G Ex ec nC IIC T4 Gc)
  Ex ec nC (increased safety, enclosed break)
  IP 65 Housing EL3040
  EL3000 analyzer modules
  Uras26
  Caldos27
  Magnos28
  Combination of analyzer modules possible in one housing
  Purgeable analyzer housing for corrosion protection
  Non flammable sample gas
  Certified in accordance to:
  in preparation
Cat. 3G / EPL Gc, non flammable gases
EL3000
Zone 2
Zone 1
Zone 0
Purgeable
Explosion proof
◄
Back to previous section

## Slide 113: Zone 1
EL3060
Cat. 2G / EPL Gb
  Explosion proof design: Ex db eb IIC T4 Gb
  Explosion proof design: Ex db IIC T4 Gb
  IIC classification, able for Acetylene-Hydrogen ambiances
  Flammable or non flammable sample gas (Zone 1 gas)
  Analyzer can be purged for corrosion protection
  Uras26, Magnos28, Caldos25/27
  Combi-Analyzer System
  Certified in accordance to:
Zone 2
Zone 1
Zone 0
Zone 1 gas
Purgeable
Explosion proof
◄
Back to previous section

## Slide 114: General Purpose
Flammable Gases
  Manufacture Declaration
  flammable sample gas
  IP54 Housing, housing purge necessary
  AO2000 analyzer modules
  Uras26
  Caldos25/27
  Magnos27/28
  Limas21 UV
  Multi-analyzer systems possible
  Stainless steel tubing (Uras/Limas) and direct access analyzers (Caldos/Magnos)
  Flame Barriers (optional)
AO2000
Zone 2
Zone 1
Zone 0
Gas
Purged
Non-Ex
◄

## Slide 115: Zone 2
Cat. 3G / EPL Gc, non flammable gases
  Manufacture Declaration (II 3G Ex nA nC II T4 Gc)
  Ex nA nC (non incendive)
  non flammable sample gas
  IP54 Housing (AO2020 / AO2040)
  purgeable housing for corrosion protection
  AO2000 analyzer modules
  Uras26
  Caldos25/27
  Magnos27/28
  Limas21 UV
  Multi-analyzer systems possible
  Certified in accordance to:
  in preparation
AO2000
Zone 2
Zone 1
Zone 0
Gas
Purgeable
Explosion proof
◄

## Slide 116: Zone 2
Cat. 3G / EPL Gc, flammable gases
  Manufacture Declaration (II 3G Ex nA pyb II T4 Gc)
  Ex nA (non incendive)
  Ex pyb (purged and controlled sensor units)
  Flammable / non flammable sample gas
  IP54 Housing (AO2020 / AO2040)
  purgeable housing for corrosion protection
  AO2000 analyzer modules
  Uras26
  Caldos25/27
  Magnos27/28
  Limas21 UV
  Multi-analyzer systems possible
  Certified in accordance to:
  in preparation
AO2000
Zone 2
Zone 1
Zone 0
Gas
Purgeable
Explosion proof
◄

## Slide 117: Class I, Division 2
CSA Approval
  Standard system approved for Class I, Div. 2, Groups A B C D, Temp. Class T4
  Measuring of flammable gases and non-flammable gases
  Analyzers:
  Uras26
  Magnos27/28
  Caldos25/27
  Limas11/21 UV
  LS25
  Sampling conditioning:
  SCC-C
  SCC-F
AO2000
Back to previous section
◄

## Slide 118: Safety Integrity Level (SIL)
  SIL Manufacturer Declaration
    Specific to Magnos28 paramagnetic O2 analyzer
    Acc. to EN 61508 (2010) part 2
    SIL assessment carried out by independent consultant (embeX)
  Available housings
    EL2020 / EL3040 / EL3060
  Possible SIL Levels
    SIL 2 via single use (One analyzer system)
    SIL 3 via dual use (Redundant analyzer system)
EasyLine analyzer with SIL declaration
EL3000 / EL3060 with Magnos28
◄
Back to previous section